## Dilutor 1
date: 18.06.24 | [13 empty] | SWAVE
before: rinsing of tubes with 3 mL inside and outside, in between rinsing after each aspiration (2 mL, 10 mL/min) | [13 empty] | Dilutor 1
Conditions:
the needle tip was inside the dispensed volume at the end of the dispensing step --> no drops were taken up with the needle tip
number of each volume | wanted volume | airgap volume | extra volume | draw speed (aspiration and dispensing) | equilibration time (aspiration and dispensing =) | multiple aspirations | liquid | rinsing volume (after each transfer (inside and outside volume =) | sample name | used dilutor (1 mL) | mass of vial | measured mass | mass of sample | calculated volume (density of water = 1 g/mL) |  | deviation | deviation
| [mL] | [µL] | [mL] | [ml/min] | [s] |  |  | [mL] |  | 1 | [g] | [g] | [g] | [mL] |  | [µL] | [% of wanted volume]
1 | 0.2 | 50 | 0.05 | 1 | 2 | no | water | 2 | Dil1-0.2-1-w | 1 | 2.5911 | 2.7881 | 0.197 | 0.197 |  | -3 | -1.5%
2 | 0.2 | [7 empty] | Dil1-0.2-2-w | 1 | 2.6024 | 2.8001 | 0.1977 | 0.1977 |  | -2.3 | -1.15%
3 | 0.2 | [7 empty] | Dil1-0.2-3-w | 1 | 2.6121 | 2.8088 | 0.1967 | 0.1967 |  | -3.3 | -1.65%
1 | 0.3 | 50 | 0.05 | 1 | 2 | no | water | 2 | Dil1-0.3-1-w | 1 | 2.6095 | 2.9051 | 0.29557 | 0.29557 |  | -4.43 | -1.477%
2 | 0.3 | [7 empty] | Dil1-0.3-2-w | 1 | 2.6251 | 2.921 | 0.29592 | 0.29592 |  | -4.08 | -1.36%
3 | 0.3 | [7 empty] | Dil1-0.3-3-w | 1 | 2.6096 | 2.9054 | 0.29582 | 0.29582 |  | -4.18 | -1.393%
1 | 0.4 | 50 | 0.05 | 1 | 2 | no | water | 2 | Dil1-0.4-1-w | 1 | 2.5802 | 2.9765 | 0.3963 | 0.3963 |  | -3.7 | -0.925%
2 | 0.4 | [7 empty] | Dil1-0.4-2-w | 1 | 2.5942 | 2.9898 | 0.39562 | 0.39562 |  | -4.38 | -1.095%
3 | 0.4 | [7 empty] | Dil1-0.4-3-w | 1 | 2.562 | 2.9577 | 0.3957 | 0.3957 |  | -4.3 | -1.075%
1 | 0.5 | 50 | 0.05 | 1 | 2 | no | water | 2 | Dil1-0.5-1-w | 1 | 2.5974 | 3.0938 | 0.4964 | 0.4964 |  | -3.6 | -0.72%
2 | 0.5 | [7 empty] | Dil1-0.5-2-w | 1 | 2.6394 | 3.1356 | 0.49618 | 0.49618 |  | -3.82 | -0.764%
3 | 0.5 | [7 empty] | Dil1-0.5-3-w | 1 | 2.6353 | 3.1313 | 0.49603 | 0.49603 |  | -3.97 | -0.794%
1 | 0.6 | 50 | 0.05 | 1 | 2 | no | water | 2 | Dil1-0.6-1-w | 1 | 2.5987 | 3.1947 | 0.59599 | 0.59599 |  | -4.01 | -0.6683%
2 | 0.6 | [7 empty] | Dil1-0.6-2-w | 1 | 2.5929 | 3.1891 | 0.59624 | 0.59624 |  | -3.76 | -0.6267%
3 | 0.6 | [7 empty] | Dil1-0.6-3-w | 1 | 2.6253 | 3.2213 | 0.59596 | 0.59596 |  | -4.04 | -0.6733%
1 | 0.7 | 50 | 0.05 | 1 | 2 | no | water | 2 | Dil1-0.7-1-w | 1 | 2.6193 | 3.3145 | 0.6952 | 0.6952 |  | -4.8 | -0.6857%
2 | 0.7 | [7 empty] | Dil1-0.7-2-w | 1 | 2.5904 | 3.2855 | 0.69506 | 0.69506 |  | -4.94 | -0.7057%
3 | 0.7 | [7 empty] | Dil1-0.7-3-w | 1 | 2.5874 | 3.2827 | 0.69527 | 0.69527 |  | -4.73 | -0.6757%
1 | 0.85 | 50 | 0.05 | 1 | 2 | no | water | 2 | Dil1-0.85-1-w | 1 | 2.5917 | 3.4362 | 0.84451 | 0.84451 |  | -5.49 | -0.6459%
2 | 0.85 | [7 empty] | Dil1-0.85-2-w | 1 | 2.5957 | 3.4406 | 0.84488 | 0.84488 |  | -5.12 | -0.6024%
3 | 0.85 | [7 empty] | Dil1-0.85-3-w | 1 | 2.5781 | 3.4227 | 0.84464 | 0.84464 |  | -5.36 | -0.6306%
1 | 1 | 50 | 0.05 | 1 | 2 | no | water | 2 | Dil1-1-1-w | 1 | 2.6366 | 3.6309 | 0.99428 | 0.99428 | two aspirations necessary (first 0.9, then 0.1 mL9 | -5.72 | -0.572%
2 | 1 | [7 empty] | Dil1-1-2-w | 1 | 2.6162 | 3.6114 | 0.99522 | 0.99522 | two aspirations necessary (first 0.9, then 0.1 mL9 | -4.78 | -0.478%
3 | 1 | [7 empty] | Dil1-1-3-w | 1 | 2.5779 | 3.5729 | 0.99505 | 0.99505 | two aspirations necessary (first 0.9, then 0.1 mL9 | -4.95 | -0.495%

## Dilutor 2
date: 18.06.24 | [13 empty] | SWAVE
before: rinsing of tubes with 3 mL inside and outside, in between rinsing after each aspiration | [13 empty] | Dilutor 2
Conditions:
the needle tip was inside the dispensed volume at the end of the dispensing step --> no drops were taken up with the needle tip
number of each volume | wanted volume | airgap volume | extra volume | draw speed (aspiration and dispensing) | equilibration time (aspiration and dispensing =) | multiple aspirations | liquid | rinsing volume (after each transfer (inside and outside volume =) | sample name | used dilutor (1 mL) | mass of vial | measured mass | mass of sample | calculated volume (density of water = 1 g/mL) |  | deviation | deviation
| [mL] | [µL] | [mL] | [ml/min] | [s] |  |  | [mL] |  | 2 | [g] | [g] | [g] | [mL] |  | [µL] | [% of wanted volume]
1 | 0.2 | 50 | 0.05 | 1 | 2 | no | water | 2 | Dil2-0.2-1-w | 2 | 2.6518 | 2.8483 | 0.1965 | 0.1965 |  | -3.5 | -1.75%
2 | 0.2 | [7 empty] | Dil2-0.2-2-w | 2 | 2.6519 | 2.8487 | 0.1968 | 0.1968 |  | -3.2 | -1.6%
3 | 0.2 | [7 empty] | Dil2-0.2-3-w | 2 | 2.66 | 2.857 | 0.197 | 0.197 |  | -3 | -1.5%
1 | 0.3 | 50 | 0.05 | 1 | 2 | no | water | 2 | Dil2-0.3-1-w | 2 | 2.6436 | 2.9387 | 0.2951 | 0.2951 |  | -4.9 | -1.633%
2 | 0.3 | [7 empty] | Dil2-0.3-2-w | 2 | 2.6627 | 2.9581 | 0.2954 | 0.2954 |  | -4.6 | -1.533%
3 | 0.3 | [7 empty] | Dil2-0.3-3-w | 2 | 2.6101 | 2.906 | 0.2959 | 0.2959 |  | -4.1 | -1.367%
1 | 0.4 | 50 | 0.05 | 1 | 2 | no | water | 2 | Dil2-0.4-1-w | 2 | 2.6428 | 3.0379 | 0.3951 | 0.3951 |  | -4.9 | -1.225%
2 | 0.4 | [7 empty] | Dil2-0.4-2-w | 2 | 2.6113 | 3.0067 | 0.3954 | 0.3954 |  | -4.6 | -1.15%
3 | 0.4 | [7 empty] | Dil2-0.4-3-w | 2 | 2.6009 | 2.9964 | 0.3955 | 0.3955 |  | -4.5 | -1.125%
1 | 0.5 | 50 | 0.05 | 1 | 2 | no | water | 2 | Dil2-0.5-1-w | 2 | 2.5906 | 3.0864 | 0.4958 | 0.4958 |  | -4.2 | -0.84%
2 | 0.5 | [7 empty] | Dil2-0.5-2-w | 2 | 2.5766 | 3.0725 | 0.4959 | 0.4959 |  | -4.1 | -0.82%
3 | 0.5 | [7 empty] | Dil2-0.5-3-w | 2 | 2.5911 | 3.0867 | 0.4956 | 0.4956 |  | -4.4 | -0.88%
[13 empty] | 0 | 0
1 | 0.6 | 50 | 0.05 | 1 | 2 | no | water | 2 | Dil2-0.6-1-w | 2 | 2.5935 | 3.1892 | 0.5957 | 0.5957 |  | -4.3 | -0.7167%
2 | 0.6 | [7 empty] | Dil2-0.6-2-w | 2 | 2.5813 | 3.1766 | 0.5953 | 0.5953 |  | -4.7 | -0.7833%
3 | 0.6 | [7 empty] | Dil2-0.6-3-w | 2 | 2.6461 | 3.242 | 0.5959 | 0.5959 |  | -4.1 | -0.6833%
1 | 0.7 | 50 | 0.05 | 1 | 2 | no | water | 2 | Dil2-0.7-1-w | 2 | 2.634 | 3.3285 | 0.6945 | 0.6945 |  | -5.5 | -0.7857%
2 | 0.7 | [7 empty] | Dil2-0.7-2-w | 2 | 2.5832 | 3.2784 | 0.6952 | 0.6952 |  | -4.8 | -0.6857%
3 | 0.7 | [7 empty] | Dil2-0.7-3-w | 2 | 2.6602 | 3.355 | 0.6948 | 0.6948 |  | -5.2 | -0.7429%
1 | 0.85 | 50 | 0.05 | 1 | 2 | no | water | 2 | Dil2-0.85-1-w | 2 | 2.5887 | 3.4325 | 0.8438 | 0.8438 |  | -6.2 | -0.7294%
2 | 0.85 | [7 empty] | Dil2-0.85-2-w | 2 | 2.6268 | 3.47 | 0.8432 | 0.8432 |  | -6.8 | -0.8%
3 | 0.85 | [7 empty] | Dil2-0.85-3-w | 2 | 2.6379 | 3.4816 | 0.8437 | 0.8437 |  | -6.3 | -0.7412%
1 | 1 | 50 | 0.05 | 1 | 2 | no | water | 2 | Dil2-1-1-w | 2 | 2.3733 | 3.3654 | 0.99213 | 0.99213 | two aspirations necessary (first 0.9, then 0.1 mL9 | -7.87 | -0.787%
2 | 1 | [7 empty] | Dil2-1-2-w | 2 | 2.3899 | 3.3824 | 0.99248 | 0.99248 | two aspirations necessary (first 0.9, then 0.1 mL9 | -7.52 | -0.752%
3 | 1 | [7 empty] | Dil2-1-3-w | 2 | 2.323 | 3.3153 | 0.99227 | 0.99227 | two aspirations necessary (first 0.9, then 0.1 mL9 | -7.73 | -0.773%

## Dilutor 3
date: 18.06.24 | [13 empty] | SWAVE
before: rinsing of tubes with 10 mL inside and outside, in between rinsing after each aspiration (e.g. three aspirations for 3 mL) | [13 empty] | Dilutor 3
Conditions:
the needle tip was inside the dispensed volume at the end of the dispensing step --> no drops were taken up with the needle tip
number of each volume | wanted volume | airgap volume | extra volume | draw speed (aspiration and dispensing) | equilibration time (aspiration and dispensing =) | multiple aspirations | liquid | rinsing volume (after each transfer (inside and outside volume =) | sample name | used dilutor (10 mL) | mass of vial | measured mass | mass of sample | calculated volume (density of water = 1 g/mL) | deviation | deviation
| [mL] | [µL] | [mL] | [ml/min] | [s] |  |  | [mL] |  | 3 | [g] | [g] | [g] | [mL] | [µL] | [% of wanted volume]
1 | 0.5 | 50 | 0.05 | 1 | 2 | no | water | 2 | Dil3-0.5-1-w | 3 | 2.5977 | 3.0892 | 0.4915 | 0.4915 | -8.5 | -1.7%
2 | 0.5 | [7 empty] | Dil3-0.5-2-w | 3 | 2.603 | 3.0952 | 0.4922 | 0.4922 | -7.8 | -1.56%
3 | 0.5 | [7 empty] | Dil3-0.5-3-w | 3 | 2.6101 | 3.1035 | 0.4934 | 0.4934 | -6.6 | -1.32%
1 | 0.6 | 50 | 0.05 | 1 | 2 | no | water | 2 | Dil3-0.6-1-w | 3 | 2.592 | 3.1854 | 0.5934 | 0.5934 | -6.6 | -1.1%
2 | 0.6 | [7 empty] | Dil3-0.6-2-w | 3 | 2.589 | 3.1829 | 0.5939 | 0.5939 | -6.1 | -1.017%
3 | 0.6 | [7 empty] | Dil3-0.6-3-w | 3 | 2.5851 | 3.1797 | 0.5946 | 0.5946 | -5.4 | -0.9%
1 | 0.7 | 50 | 0.05 | 1 | 2 | no | water | 2 | Dil3-0.7-1-w | 3 | 2.6002 | 3.2915 | 0.6913 | 0.6913 | -8.7 | -1.243%
2 | 0.7 | [7 empty] | Dil3-0.7-2-w | 3 | 2.611 | 3.3016 | 0.6906 | 0.6906 | -9.4 | -1.343%
3 | 0.7 | [7 empty] | Dil3-0.7-3-w | 3 | 2.5828 | 3.2738 | 0.691 | 0.691 | -9 | -1.286%
1 | 0.85 | 50 | 0.05 | 1 | 2 | no | water | 2 | Dil3-0.85-1-w | 3 | 2.5655 | 3.4051 | 0.8396 | 0.8396 | -10.4 | -1.224%
2 | 0.85 | [7 empty] | Dil3-0.85-2-w | 3 | 2.5957 | 3.4354 | 0.8397 | 0.8397 | -10.3 | -1.212%
3 | 0.85 | [7 empty] | Dil3-0.85-3-w | 3 | 2.5702 | 3.4103 | 0.8401 | 0.8401 | -9.9 | -1.165%
1 | 1 | 50 | 0.05 | 1 | 2 | no | water | 2 | Dil3-1-1-w | 3 | 2.6275 | 3.6168 | 0.98935 | 0.98935 | -10.65 | -1.065%
2 | 1 | [7 empty] | Dil3-1-2-w | 3 | 2.5729 | 3.5624 | 0.9895 | 0.9895 | -10.5 | -1.05%
3 | 1 | [7 empty] | Dil3-1-3-w | 3 | 2.6153 | 3.6051 | 0.98985 | 0.98985 | -10.15 | -1.015%
1 | 3 | 50 | 0.05 | 3.5 | 2 | no | water | 3 | Dil3-3-1-w | 3 | 10.7757 | 13.7575 | 2.9818 | 2.9818 | -18.2 | -0.6067%
2 | 3 | [7 empty] | Dil3-3-2-w | 3 | 10.7655 | 13.7479 | 2.9824 | 2.9824 | -17.6 | -0.5867%
3 | 3 | [7 empty] | Dil3-3-3-w | 3 | 10.8262 | 13.809 | 2.9828 | 2.9828 | -17.2 | -0.5733%
1 | 3.5 | 50 | 0.05 | 3.5 |  | no | water | 3.5 | Dil3-3.5-1-w | 3 | 10.7039 | 14.1813 | 3.4774 | 3.4774 | -22.58 | -0.6451%
2 | 3.5 | [7 empty] | Dil3-3.5-2-w | 3 | 10.8421 | 14.3194 | 3.4773 | 3.4773 | -22.7 | -0.6486%
3 | 3.5 | [7 empty] | Dil3-3.5-3-w | 3 | 10.6641 | 14.1393 | 3.4752 | 3.4752 | -24.8 | -0.7086%
2
1 | 4 | 50 | 0.05 | 4 |  | no | water | 4 | Dil3-4-1-w | 3 | 10.8723 | 14.8486 | 3.9763 | 3.9763 | -23.7 | -0.5925%
2 | 4 | [7 empty] | Dil3-4-2-w | 3 | 10.5696 | 14.5464 | 3.9768 | 3.9768 | -23.2 | -0.58%
3 | 4 | [7 empty] | Dil3-4-3-w | 3 | 10.7259 | 14.7021 | 3.9762 | 3.9762 | -23.85 | -0.5962%
1 | 5 | 50 | 0.05 | 5 | 2 | no | water | 5 | Dil3-5-1-w | 3 | 11.2158 | 16.1869 | 4.9711 | 4.9711 | -28.9 | -0.578%
2 | 5 | [7 empty] | Dil3-5-2-w | 3 | 11.2595 | 16.2307 | 4.9712 | 4.9712 | -28.8 | -0.576%
3 | 5 | [7 empty] | Dil3-5-3-w | 3 | 11.3877 | 16.3589 | 4.9712 | 4.9712 | -28.8 | -0.576%
1 | 5.5 | 50 | 0.05 | 5.5 | 2 | no | water | 5.5 | Dil3-5.5-1-w | 3 | 11.2824 | 16.7492 | 5.4668 | 5.4668 | -33.2 | -0.6036%
2 | 5.5 | [7 empty] | Dil3-5.5-2-w | 3 | 11.5312 | 16.9995 | 5.4683 | 5.4683 | -31.7 | -0.5764%
3 | 5.5 | [7 empty] | Dil3-5.5-3-w | 3 | 11.3702 | 16.8334 | 5.4632 | 5.4632 | -36.8 | -0.6691%
1 | 6 | 50 | 0.05 | 6 | 2 | no | water | 6 | Dil3-6-1-w | 3 | 10.7122 | 16.6766 | 5.9644 | 5.9644 | -35.57 | -0.5928%
2 | 6 | [7 empty] | Dil3-6-2-w | 3 | 10.8103 | 16.7767 | 5.9664 | 5.9664 | -33.58 | -0.5597%
3 | 6 | [7 empty] | Dil3-6-3-w | 3 | 10.7032 | 16.6676 | 5.9644 | 5.9644 | -35.55 | -0.5925%
1 | 7 | 50 | 0.05 | 7 | 2 | no | water | 7 | Dil3-7-1-w | 3 | 10.7613 | 17.7175 | 6.9562 | 6.9562 | -43.83 | -0.6261%
2 | 7 | [7 empty] | Dil3-7-2-w | 3 | 10.7497 | 17.7118 | 6.9621 | 6.9621 | -37.9 | -0.5414%
3 | 7 | [7 empty] | Dil3-7-3-w | 3 | 10.7766 | 17.7396 | 6.963 | 6.963 | -36.99 | -0.5284%